## Slide 1: Steerer amplifier load characterization
What the EEL5000 amplifiers drive, how we measure it, and where the limits are
Isaac Legault · Hirst Lab · October 2026
Notes: This is about the four high-voltage amplifiers that drive the steerer plates. The goal is to know exactly what each amplifier is driving and how close we run to its limits, because we want to run them for long irradiations without them ever hitting LIMIT or TRIP.

## Slide 2: One chain per steerer plate
Function generator
DG1022Z
±5 V
HV amplifier
EEL5000
×1000, up to ±5 kV
HV cable
~6 ft coax
Feedthrough
into vacuum
Steerer plate
X+, X-, Y+ or Y-
Built-in monitors
Voltage: 1 V = 1 kV
Current: 1 V = 10 mA
LabJack T7
records both monitors in the RBL app
Everything the app knows about the load comes from those two monitors.
Notes: There are four of these chains, one for each plate. The function generator makes a small waveform, the amplifier multiplies it by 1000, and the cable carries it into the vacuum to the plate. Each amplifier has two monitor outputs, one for voltage and one for current, and the LabJack records both. That is the only window we have into what the amplifier is actually doing.

## Slide 3: Current only flows while the voltage is changing
I = C · dV/dt
The plates, cable and amplifier output act as a capacitor C.
A triangle wave changes voltage at a constant rate, so it draws a constant current that flips sign at each corner.
Steady current: I = 4 · f · C · V
Example: 500 Hz, ±2 kV, C = 1.6 nF
Chart: 
Chart: Time (ms)
| 0 | 0.5 | 0.5 | 1.5 | 1.5 | 2.5 | 2.5 | 3.5 | 3.5 | 4
Current | 6.4 | 6.4 | -6.4 | -6.4 | 6.4 | 6.4 | -6.4 | -6.4 | 6.4 | 6.4
Notes: This one equation is the whole model. A capacitor only takes current while you are changing its voltage. Hold a voltage steady and almost no current flows, only leakage. Because the raster uses a triangle wave, the voltage changes at a constant rate, so the current is flat and just reverses at each turn. In this example that is about 6.4 mA. Double the frequency or the amplitude and the current doubles. So if we know C, we can predict the current before we ever turn the amplifier on. That is why measuring C matters.

## Slide 4: Three independent ways to measure C
A
Impedance sweep
Sine waves at 9 frequencies.
Compares the timing of current and voltage. A pure capacitor leads by exactly 90°; anything less is leakage.
B
DC leakage
Holds steady voltages for 10 s each.
Averaging removes the monitor noise, so leakage down to about 1 µA is visible.
C
Charge per edge
Square wave steps, 0.5 to 5 kV.
Adds up the charge in each edge: Q = C · ΔV. Unaffected by the monitor's limited speed.
A and C use different physics. If they give the same C, the load really is a capacitor.
Notes: We do not trust a single number. Method A drives sine waves and looks at the timing between current and voltage: a perfect capacitor's current is exactly a quarter cycle ahead, and any shortfall is leakage. Method B holds DC and averages for ten seconds, which pulls a microamp-level leak out of the noise. Method C uses square-wave steps and adds up the charge in each spike. Charge is preserved even when the monitor smooths the spike, so this one does not depend on how fast the monitor is. If A and C agree, we have a real capacitance and not an artifact.

## Slide 5: Where the capacitance lives
What is connected | X+ | X- | Y+ | Y-
Amplifier alone | ____ pF | ____ pF | ____ pF | ____ pF
+ HV cable, far end open | ____ pF | ____ pF | ____ pF | ____ pF
+ feedthrough and plates | ____ pF | ____ pF | ____ pF | ____ pF
Rough estimate from geometry
Cable ~110 pF
Feedthrough ~10 pF
Plates ~3 pF
Total ~125 pF
Each row adds one piece. The step between rows is that piece's capacitance. Anything well above the estimate is inside the amplifier or the cable.
Plates: C = ε₀A/d with 12.5 × 10.2 cm plates, 3.8 cm gap. Cable: typical coax ~60 pF/m.
Notes: Measuring the total only tells us how big the load is, not where it comes from. So we build it up one piece at a time: the amplifier alone, then with the cable attached, then with everything. The difference between rows is each piece. From geometry we would expect only about 125 pF in total, and the plates themselves are only about 3 pF. Our earlier measurements came out far higher than that, which is why this breakdown matters: if most of it is inside the amplifier, we cannot reduce it; if it is in the cable, we can. That decides how fast we can raster.

## Slide 6: What each amplifier can supply
20 mA
Continuous
Any current the amplifier holds for a long time. A triangle wave's steady current must stay under this.
100 mA
Burst, at most 4 ms
Headroom for short spikes from voltage steps and arcs. Never an operating point.
100 ms
Recovery after a burst
The manual says output drops to 10 mA for this long before another burst is available.
Goal: run long irradiations with steady current well under 20 mA and spikes that never reach LIMIT or TRIP.
Source: EEL5000.20.100 manual, pp. 1-3 and 2-1. Questions on LIMIT and TRIP behaviour sent to the manufacturer.
Notes: These are the three numbers from the manual that define the safe zone. Twenty milliamps is what the amplifier can hold indefinitely. One hundred milliamps is allowed only for a few milliseconds, which is enough for the short spike you get when the voltage jumps. After a burst like that the amplifier recovers at a reduced current for a tenth of a second. I have asked the manufacturer exactly how the LIMIT and TRIP circuits decide when to act, because during an irradiation the point is that they never act.

## Slide 7: Steady current vs. raster frequency
Chart: Raster frequency (Hz)
Sloped lines
Predicted steady current, I = 4 f C V, for one amplitude each.
20 mA line
The real limit for a raster. Operate below it with margin.
100 mA line
Only for spikes lasting 4 ms or less. A raster can never run here.
Drawn for C = 1.6 nF as an illustration. The app draws each plate from its own measured C.
Notes: This is the chart that matters most for planning. Each sloped line is one amplitude. As frequency goes up, current goes up in proportion. Where a sloped line crosses 20 mA is the fastest we can raster at that amplitude. The 100 mA line is there only to show the spike headroom. A raster draws steady current, so it is held to the 20 mA line, not the 100. In the app, each plate gets its own line from its own measured capacitance, with the date it was measured.

## Slide 8: What our instruments can and cannot see
Chart: Time (µs)
LabJack samples (every 133 µs) | 6 | 6 | 6.11 | 6 | 6
Example: a 20 µs, 40 mA spike
Reliable
Steady current and capacitance
Leakage down to about 1 µA
Total charge in a spike
Not reliable
Peak of any spike shorter than about 100 µs: the monitor smooths it and samples can miss it
Those need an oscilloscope on the monitor outputs
Notes: The grey line is a short spike as it really happens. The current monitor can only respond so fast, so it shows a rounded, lower peak, about 31.9 mA instead of 40. During an irradiation the LabJack reads each monitor only every 133 microseconds, so in this example it never lands on the spike at all. That is why I only trust spike charge and steady values from the app. For the height of short spikes I will put a scope on the monitor outputs.

## Slide 9: Next tests
1
Clamp test
Raise frequency at fixed voltage until LIMIT engages. Shows where the amplifier really clamps.
2
Measure C, all four plates
Methods A and C, in all three connection setups. Fills in the capacitance table.
3
Voltage ladder
Method C from 0.5 to 5 kV. C should not change with voltage; if it does, something is discharging.
4
Swap X and Y amplifiers
Repeat the measurements. Shows whether a problem follows the amplifier or the plate.
5
12-hour run
The exact triangle wave of a real experiment, logging every current spike.
Notes: This is the plan going forward. First, find out where the amplifier really clamps instead of trusting the manual. Then measure every plate's capacitance two independent ways and break it down by what is connected. Then step the voltage up to full scale to make sure nothing changes with voltage, which would mean a discharge. Then swap the X and Y amplifiers, so we can tell whether a bad channel is the amplifier or the wiring. And finally, run the exact experimental waveform for twelve hours while logging every spike, so we have evidence that long irradiations are safe.