## Slide 1
LOAD CHARACTERIZATION · MODE A
Measuring the Plate Load
How we get the capacitance of each deflection channel, and what that capacitance sets for the safe drive limits.
Right Beam Line · Ion Beam Laboratory · UW–Madison NEEP
Notes: Two halves to this talk. First: how the app measures the capacitance of a deflection channel. Second: why that number is the thing that sets our voltage and frequency limits. No new physics here - it is all one equation, applied carefully.

## Slide 2
THE PROBLEM
Why we need this number
1
The plates are a capacitor.
Each deflection channel drives the plates plus the cable and the feedthrough. Electrically that is one capacitor, and we call its value C.
2
A capacitor's current grows with frequency and voltage.
Rastering means driving that capacitor with an AC waveform. The faster and higher we drive it, the more current the amplifier has to supply.
3
Above 20 mA the amplifier trips.
So C is what decides how high and how fast we are allowed to drive - the operating envelope is a direct consequence of it.
WHAT WE HAVE TODAY
1500 pF
One global number, back-solved once from two AC sweeps. It is used for all four channels.
It cannot tell the four channels apart.
It cannot say whether a channel is leaking.
It has never been re-measured.
Notes: The point of the Load Characterization tab is to replace that single 1500 pF constant with a real, per-channel, on-demand measurement - and to get a leakage number out of the same measurement for free.

## Slide 3
NOTATION
The symbols, in plain terms
f
Frequency (Hz)
How many times per second we swing the plate voltage.
ω
Angular frequency (rad/s)
The same thing in radians: ω = 2πf. It only appears to keep the formulas short.
V
Peak plate voltage (kV)
The height of the voltage waveform we command.
I
Peak current (mA)
The height of the current waveform the amplifier actually delivers.
Δφ
Phase angle (degrees)
How far ahead of the voltage the current waveform sits in time. This is the measurement.
C
Capacitance (pF)
How much charge the plates hold per volt applied. The number we are after.
G
Conductance (µS)
Leakage: current that flows in step with the voltage instead of ahead of it.
δ
Loss tangent, tan δ (unitless)
Leakage expressed as a fraction of the capacitive current. Zero for a perfect capacitor.
Notes: Only three of these are measured: V, I and the phase angle. C and G are computed from those three. tan delta is just a ratio of the two computed numbers.

## Slide 4
MODE A · ONE CHANNEL AT A TIME
How the capacitance is measured
1
Drive one channel with a sine wave at a known frequency.
The frequency is commanded, so we know it exactly - we are not searching for it.
2
Record both monitor BNCs at 50,000 samples per second for 1 second.
Voltage monitor and current monitor, sampled at the same instant on the same card.
3
Extract three numbers from that record.
The peak voltage V, the peak current I, and the phase angle Δφ between them.
4
Solve for C and G.
Two equations, two unknowns. Then repeat at nine frequencies from 200 Hz to 3 kHz.
THE TWO EQUATIONS
C = I · sin(Δφ) / (2πf · V)
G = I · cos(Δφ) / V
Everything on the right is measured.
Why nine frequencies?
A real capacitor should give the same C at every frequency. Nine points is how we check that it does.
Notes: The drive amplitude is not constant across the nine frequencies. It is set each time so the load draws about 4 mA - 2.65 kV at 200 Hz, down to about 180 V at 3 kHz. That keeps the signal the same distance above the noise at every point in the sweep.

## Slide 5
WHY THE PHASE MATTERS
Everything comes from one angle
Voltage
Current
Δφ
↑ this part is C
this part is G (leakage)
Angle exaggerated for clarity - a healthy channel sits within a degree of 90°.
A perfect capacitor puts the current exactly 90° ahead of the voltage.
Charge is proportional to voltage, so current - the rate of change of charge - peaks when the voltage is changing fastest, which is a quarter cycle earlier.
The straight-up part of the current gives C.
That is the sin(Δφ) term. Near 90° the sine is essentially flat, so a couple of degrees of error costs a fraction of a percent on C. C is robust.
The sideways part gives the leakage.
That is the cos(Δφ) term, and near 90° the cosine is changing fastest. A 4° error invents leakage that is not there. Leakage is fragile.
Notes: This is the single most important slide for judging the results. The capacitance is a projection onto the vertical axis and is insensitive to small angle errors. The leakage is a projection onto the horizontal axis and is maximally sensitive to them. So we quote C with confidence and treat the leakage as a comparison between channels, not as an absolute.

## Slide 6
THE PAYOFF
What the capacitance sets: the drive limit
CURRENT DRAWN
I = k · f · C · V
k = 2π for a sine, 4 for a triangle
HIGHEST VOLTAGE ALLOWED
V = 20 mA / (k · f · C)
the same equation, solved for V at the trip current
Get C wrong and this limit is wrong. That is the entire reason the measurement exists.
Chart: Highest safe peak voltage vs frequency (C = 1200 pF, 20 mA trip)
| 200 | 400 | 600 | 800 | 1000 | 1200 | 1400 | 1600 | 1800 | 2000 | 2200 | 2400 | 2600 | 2800 | 3000
Sine drive | 5 | 5 | 4.421 | 3.316 | 2.653 | 2.21 | 1.895 | 1.658 | 1.474 | 1.326 | 1.206 | 1.105 | 1.02 | 0.947 | 0.884
Triangle drive | 5 | 5 | 5 | 5 | 4.167 | 3.472 | 2.976 | 2.604 | 2.315 | 2.083 | 1.894 | 1.736 | 1.603 | 1.488 | 1.389
Flat section = the amplifier's 5 kV rating is the limit. Falling section = the load current is the limit.
Notes: Read the curve from the left. Below about 530 Hz for a sine the amplifier's own 5 kV rating is what stops us. Above that the load current takes over and the ceiling falls off as one over frequency. A triangle draws less peak current at the same amplitude, so its crossover sits higher, near 830 Hz. Both curves scale directly with C - a load 20 percent larger pushes the whole falling section down by 20 percent.

## Slide 7
WHERE THIS LEAVES US
What to trust, and what to do next
CAPACITANCE
Trustworthy.
Good to a few percent, per channel, on demand. This is what should replace the single 1500 pF constant in the drive limits.
LEAKAGE
Comparative only.
It is a phase measurement, and the monitors' own phase response contaminates it above roughly 1 kHz. A channel that reads high against the other three is real; the absolute number is not.
NEXT STEP
Eight sweeps.
All four channels, disconnected and then on the plates. The difference between the two is the capacitance of the plates themselves.
One sine, one angle, one equation — and the drive envelope follows from it.
Notes: Ask for a decision on one thing: whether to run the eight sweeps now, and whether to buy a known-good reference capacitor so the leakage number can be calibrated rather than only compared.