## Slide 1
The symbols
f
Frequency (Hz)
How fast we swing the plate voltage.
V
Peak voltage (kV)
Height of the voltage waveform.
I
Peak current (mA)
Height of the current waveform.
Δφ
Phase angle (degrees)
How far ahead of the voltage the current sits.
C
Capacitance (pF)
Charge the plates hold per volt. What we want.
G
Conductance (µS)
Leakage current.
We measure V, I and Δφ. C and G are calculated from those three.
Notes: Only three things are measured. Everything else is arithmetic.

## Slide 2
How we measure the capacitance
1.
Drive one channel with a sine wave.
At a frequency we choose, so we know it exactly.
2.
Record the voltage and current monitors.
50,000 samples per second, for one second.
3.
Pull out three numbers.
Peak voltage V, peak current I, and the angle Δφ between them.
4.
Solve for C and G.
Then repeat at nine frequencies, 200 Hz to 3 kHz.
C = I · sin(Δφ) / (2πf · V)
G = I · cos(Δφ) / V
A real capacitor gives the same C at every frequency. Nine points is how we check that it does.
Notes: The drive voltage is not the same at every frequency. It is set so the load draws about 4 mA each time - 2.65 kV at 200 Hz, about 180 V at 3 kHz.

## Slide 3
It all comes from one angle
Voltage
Current
Δφ
this part gives C
this part gives G
Angle exaggerated. A good channel sits within a degree of 90°.
A perfect capacitor puts the current 90° ahead of the voltage.
Current is the rate of change of charge, so it peaks a quarter cycle before the voltage does.
The upright part gives C.
Near 90° this barely changes, so a couple of degrees of error costs a fraction of a percent. C is solid.
The sideways part gives the leakage.
Near 90° this changes fast. A few degrees of error invents leakage that is not there.
Notes: Capacitance is the vertical leg and is insensitive to small angle errors. Leakage is the horizontal leg and is very sensitive to them. So we quote C with confidence and use leakage only to compare channels against each other.

## Slide 4
What C sets: the drive limit
Current the amplifier has to supply
I = k · f · C · V
k = 2π for a sine, 4 for a triangle
Highest voltage allowed, at the 20 mA trip
V = 20 mA / (k · f · C)
the same equation, rearranged
Get C wrong and this limit is wrong.
Chart: Highest safe peak voltage (C = 1200 pF, 20 mA trip)
| 200 | 400 | 600 | 800 | 1000 | 1200 | 1400 | 1600 | 1800 | 2000 | 2200 | 2400 | 2600 | 2800 | 3000
Sine | 5 | 5 | 4.421 | 3.316 | 2.653 | 2.21 | 1.895 | 1.658 | 1.474 | 1.326 | 1.206 | 1.105 | 1.02 | 0.947 | 0.884
Triangle | 5 | 5 | 5 | 5 | 4.167 | 3.472 | 2.976 | 2.604 | 2.315 | 2.083 | 1.894 | 1.736 | 1.603 | 1.488 | 1.389
Flat = the amplifier's 5 kV rating is the limit. Falling = the load current is the limit.
Notes: Below about 530 Hz for a sine, the amplifier's own 5 kV rating is what stops us. Above that the load current takes over and the ceiling falls off as one over frequency. A triangle draws less peak current at the same amplitude, so its crossover is higher, near 830 Hz. Both curves scale directly with C.